How to acquire samples using Prologix and HP53131A

Connect Prologix to computer USB
Check what is the assigned com port
Use Hyperterminal (or equivalent) to open the port at 19200 8bits, no parity, RTS-CTS

You can test if Prologix work by sending the following command

++help

And you should see a list of available commands

# Setting the GPIB address

To find out or set the instrument address you need to enter Utility menu

Hold the Recall Key and cycle the Power key

You should now be in the Utility menu and can use the Recall key to move to the HP-IB: subpart, then use the arrows and then the Enter key to actually set the address

It is assumed that the instrument address will be 3

then you can use the following Prologix command to set the address

++addr 3

Note that there is no echo and that to check the result you can use the command

++addr

to show currently configured address, the parameter is automatically saved.

# Verify the instrument actually respond

Type the following command

*idn?

You should get the following response

HEWLETT-PACKARD,53131A,0,3703

# Send a command to instrument
The avilable commands depend on the instrument, so you should read the HP manual
In this case what I want is to read a frequency, so use

:read:freq?

You should receive something like

+3.276815268E+004